## Slide 1: Low level measurements of environment samples with HPGe detectors at LSM
Anne de Vismes Ott, Sébastien Aubry, Séverine Bonnot, Neila Mohara, Christophe Ardois

## Slide 2: Introduction
Institute for Radiation Protection and Nuclear Safety
Laboratory of Environmental Radioactivity Metrology
  Activity determination in environmental samples
    Environment surveillance
    Radioecology studies
    Emergency preparedness
  Research and development : materials and methods improvements
  Member of the IAEA Network of Analytical Laboratories for the Measurement of Environmental Radioactivity (ALMERA) and the CELLAR collaboration
Gamma ray spectrometry
  Accreditation since 1999
  1500-2000 measurements/year

## Slide 3: Introduction
Specificities
  wide range of measurements :
    Samples (solid): fauna, flora, waters, soils, sediments, aerosol filters…
    Radionuclides:
      Naturally occurring radionuclides: cosmogenic (7Be, 22Na...) and telluric (40K, U & Th natural decay series)
      Artificial radionuclides
        present at trace levels in the French environment: 60Co, 137Cs, 129I…
        potentially released in case of incident or accident: 131I, 134Cs, 106Rh…
    Counting geometries depending on the sample availability
    Detectors: coaxial, planar, semi-planar, well-type…
  low level radioactivity
         Bq  mBq / sample

## Slide 4: 22 High Purity Germanium detectors
Shallow underground laboratory in Orsay
  10 m w.e.: 2d basement, 3 m slab of borated concrete
  Shielded room: 10 cm Pb + 5 mm Cu
  12 low background:
  6 BEGe5030 + 1 BEGe6530 (Canberra) + 5 Profile-FX (Ortec)
  High relative efficiency > 50%
  Good resolution: 0.6 keV @ 46 keV ; 1.7 keV @1460 keV
    6 with anti-cosmic devices
  Anti-Compton system: 1 XtRa (Canberra) + NaI(Tl)
  Multi-detector Léda: 2 BEGe5030 (Canberra) + NaI(Tl)
  1 SAGe well-type detector: good resolution (Canberra)
Laboratoire Souterrain de Modane (LSM)
  Deep underground laboratory
  1700 m rock: 4800 m w.e.
  Ultra low background detectors
    2 coaxial
    2 well-type detectors of large volume (450 and 844 cm3 )
de Vismes Ott A. et al. (2013) Radioprotection, Vol. 44, n◦ 5, pp 613–618
Paradis H. et al. (2016) App. Radiat. Isot., Vol. 109, pp 487-492
Paradis H. et al. (2017) App. Radiat. Isot., Vol. 126, pp.179-184

## Slide 5: Background

## Slide 6: Ultra low background well type detectors
Low activity and small quantity
  High efficiency
  Low background
High versatility
  Detection efficiency calculated for each sample
  e = f (detector, radionuclide, material, filling height, mass)
  Monte Carlo simulation: MCNP-CP
    Extension of MCNPX (A. Berlizov, IAEA)
    Evaluated Nuclear Structure Data File
    Corrections on the detection efficiency:
      True Coincidence Summing (TCS) effect: decay scheme and measurement configuration (detector + geometry)
      Self-attenuation effect: sample material (density + composition)
    Models
      fitted with measurements of standard sources
      validated with measurements of reference materials
XXL
PPT

## Slide 7: Example
High efficiency (80%) at low energy
TCS effect : a factor of 2 for Co-60 and Y-88
e ~ 80%
FCTCS=2-2.4

## Slide 8: Atmospheric samples: aerosol filters
Aerosol filters
  OPERA-Air network
  High volume air samplers; weekly sampling
   volume ~ 100 000 m3
  Lowest 137Cs activity concentration in the air: Alençon (oceanic influence), Puy-de-Dôme and Pic-du-Midi (high altitude)
  Measurements on coaxial detectors in Modane for 2-4 days
  137Cs activity:
  0.5–100 mBq
  Detection limits :
  137Cs ~ 40 nBq.m-3
10 mL
TGD (Orsay)
aerosol
filter

## Slide 9: Atmospheric samples: aerosol filters
OPERA-Air network : multiple purpose
  Aim of radiological environment surveillance
    Baseline of the 137Cs activity concentration in the atmosphere
    Detections resulting from low-magnitude incident releases to accident releases at remote places
  In case of emergency
    Rapid and reliable information to the population
    Low level measurement to help source term assessment
  Role as a research network
    Data for transfer studies : contaminated air masses from Fukushima accident
    Improvement of the deposition calculation codes
de Vismes Ott, A. et al. (2013) J. Environ. Radioact., 125, pp. 6-16
Masson, O. et al. (2018) Environ. Sci. Technol. 2018, 52, 8488−8500
Masson, O. et al. (2019) PNAS 116 (34) 16750-16759
Fukushima accident
131I in 2017
106Rh in 2017

## Slide 10: Atmospheric samples: fallout
Cloud water samples in Puy de Dôme (1465 m)
  monthly sampling
  Volume < 1 L  dry residue mass ~ 1-100 mg
  137Cs activity < 1 mBq /sample
  137Cs activity concentration in the cloud waters: 1-5 mBq.L-1
Rain water samples in Clermont-Ferrand (645 m)
  monthly sampling; dry residue
  Volume: 20 – 300 L
  137Cs activity ~ 1 mBq /sample
  137Cs activity concentration in the rain waters: 10 - 150 µBq.L-1
Bourcier, L. et al. (2014) J. Environ. Radioact., , 128, 15-19
Masson O. et al.(2015) Atmospheric Research, 151, 45-51
Tav J. et al. (2018) Aerosol and Air Quality Research, 18: 103–113
Results
  Study of the scavenging efficiency of precipitations
    below the cloud base (i.e. washout mechanism)
    in addition to the rainout mechanism (i.e. in the cloud).
  Improved capability in modeling radionuclide deposition in case of a nuclear accident
    in foggy conditions at lowland locations
    in cloudy conditions at high altitude locations.

## Slide 11: Some times… some “exotic” samples…
Material for 3D printer
  Wire spool of PLA (polylactic acid)
  60 mL geometry
  Coaxial detector LYS
  560 000 s (~ 1 week)
  Detection efficiency transfer (Gespecor)
3D printed pieces
Results
  Detection limits (ISO 11929; a = b = 2.5 %) : 10130 mBq/kg
  3D printing is becoming a very useful option for small pieces in the sets-up of rare event searches
  The radiopurity of this particular material for 3D printing allows to use it in the alternative tower structure design of the CUPID experiment, the next generation neutrinoless double beta decay experiment.

## Slide 12: Partage
Our experience
All our 4 detectors moved in the new Partage shield in 2021

## Slide 13: 52E (Hermine)
Chart: 
| 42366.5 | 42432.5 | 42492.5 | 42548.476388888892 | 42615.5 | 42671.5 | 42727.5 | 42800.55 | 42857.5 | 42912.469386574077 | 42971.5 | 43031.444444444445 | 43090.5 | 43157.5 | 43216.5 | 43276.5 | 43343.5 | 43409.5 | 43458.5 | 43521.408333333333 | 43584.5 | 43644.5 | 43710.5 | 43769.5 | 43822.415682870371 | 43885.439583333333 | 43945.439583333333 | 44012.5 | 44074.417638888888 | 44130.397372685184 | 44189.5 | 44249.384675925925 | 44301.713865740741 | 44407.603472222225 | 44438.429861111108 | 44498.463888888888
Taux Comp Bi-214 à 609 keV | 0.69429 | 0.72 | 0.64286 | 0.56571 | 0.54 | 0.51429 | 0.55256 | 0.62791 | 0.70326 | 0.50233 | 0.65302 | 0.55883 | 0.38102 | 0.39708 | 0.49635 | 0.45495 | 0.5055 | 0.54984 | 0.61403 | 0.58845 | 0.70755 | 0.65514 | 0.51429 | 0.51429 | 0.54 | 0.43714 | 0.48857 | 0.43714 | 0.56571 | 0.54 | 0.56571 | 0.56571 | 5.2457 | 2.6229 | 5.0914 | 4.9886
Chart: 
| 42366.5 | 42432.5 | 42492.5 | 42548.476388888892 | 42615.5 | 42671.5 | 42727.5 | 42800.55 | 42857.5 | 42912.469386574077 | 42971.5 | 43031.444444444445 | 43090.5 | 43157.5 | 43216.5 | 43276.5 | 43343.5 | 43409.5 | 43458.5 | 43521.408333333333 | 43584.5 | 43644.5 | 43710.5 | 43769.5 | 43822.415682870371 | 43885.439583333333 | 43945.439583333333 | 44012.5 | 44074.417638888888 | 44130.397372685184 | 44189.5 | 44249.384675925925 | 44301.713865740741 | 44407.603472222225 | 44438.429861111108 | 44498.463888888888
Taux Comp gamme 30 keV - max | 297.2571 | 298.7743 | 298.5943 | 300.1629 | 299.9314 | 299.88 | 301.0186 | 302.4753 | 297.5023 | 295.2921 | 292.2279 | 287.7476 | 281.1686 | 270.6374 | 265.0286 | 260.281 | 257.0458 | 257.1734 | 254.8228 | 257.3045 | 257.5992 | 258.0447 | 257.4514 | 253.3371 | 253.7743 | 254.1343 | 254.4686 | 255.5486 | 256.68 | 260.0486 | 257.8114 | 260.28 | 328.1657 | 317.0571 | 350.6914 | 349.3029
Total : 250  350 d-1
Bi-214: 0.5 5 d-1
X 10
07/04/2021
No more nitrogen flushing…

## Slide 14: LYS
Chart: 
| 44347.397696759261 | 44410.436111111114 | 44469.46597222222
Taux Comp Bi-214 à 609 keV | 11.88 | 11.4171 | 16.5086
Bi-214: ~ 15 d-1
Partage

## Slide 15: XXL
Chart: 
| 42362.443865740737 | 42429.474791666667 | 42461.407638888886 | 42520.451412037037 | 42586.5 | 42654.355555555558 | 42727.420567129629 | 42765.411805555559 | 42821.474594907406 | 42881.448576388888 | 42940.427893518521 | 43006.472048611111 | 43034.476273148146 | 43073.474999999999 | 43129.478900462964 | 43189.47146990741 | 43248.475694444445 | 43307.427766203706 | 43371.469328703701 | 43437.472916666666 | 43496.397175925929 | 43553.474074074074 | 43627.659722222219 | 43675.433333333334 | 43738.403148148151 | 43802.475694444445 | 43860.472766203704 | 43931.477083333331 | 43991.45416666667 | 44043.418530092589 | 44105.462500000001 | 44323.463194444441 | 44347.4062962963 | 44420.415277777778 | 44463.458333333336
Taux Comp Bi-214 à 609 keV | 9.5236 | 5.5543 | 4.6286 | 6.0171 | 7.0971 | 5.4 | 4.185 | 4.7829 | 5.2457 | 4.7829 | 6.1714 | 4.4743 | 5.2457 | 5.0914 | 5.2457 | 3.3145 | 5.8629 | 4.9371 | 4.4743 | 3.3943 | 5.5543 | 5.5543 | 4.0114 | 4.4743 | 6.0171 | 7.56 | 4.0114 | 6.0171 | 5.0914 | 4.1657 | 15.12 | 20.3657 | 9.72 | 10.9543 | 10.8
Chart: 
| 42362.443865740737 | 42429.474791666667 | 42461.407638888886 | 42520.451412037037 | 42586.5 | 42654.355555555558 | 42727.420567129629 | 42765.411805555559 | 42821.474594907406 | 42881.448576388888 | 42940.427893518521 | 43006.472048611111 | 43034.476273148146 | 43073.474999999999 | 43129.478900462964 | 43189.47146990741 | 43248.475694444445 | 43307.427766203706 | 43371.469328703701 | 43437.472916666666 | 43496.397175925929 | 43553.474074074074 | 43627.659722222219 | 43675.433333333334 | 43738.403148148151 | 43802.475694444445 | 43860.472766203704 | 43931.477083333331 | 43991.45416666667 | 44043.418530092589 | 44105.462500000001 | 44323.463194444441 | 44347.4062962963 | 44420.415277777778 | 44463.458333333336
Taux Comp gamme 30 keV - max | 620.8036 | 567.4629 | 547.7143 | 574.0971 | 545.0914 | 572.5543 | 542.025 | 565.3029 | 562.0629 | 542.9314 | 549.8743 | 555.5829 | 579.96 | 641.8286 | 586.9029 | 571.4876 | 560.3657 | 542.6229 | 538.92 | 544.32 | 558.6686 | 893.6229 | 558.8229 | 537.5314 | 544.7829 | 636.5829 | 578.7257 | 561.1371 | 533.52 | 537.3771 | 755.5371 | 765.4114 | 725.4514 | 677.16 | 677.6229
Total : 550  700 d-1
Bi-214: 5 10 d-1
X 2
07/04/2021
Partage
No more nitrogen flushing…

## Slide 16: PPT
Chart: 
| 42366.474305555559 | 42426.53125 | 42489.461261574077 | 42552.5 | 42608.445196759261 | 42671.436793981484 | 42741.476388888892 | 42790.5 | 42857.618750000001 | 42913.404675925929 | 42969.469027777777 | 43053.459918981483 | 43368.474317129629 | 43402.401006944441 | 43461.488078703704 | 43524.442199074074 | 43584.496967592589 | 43644.461111111108 | 43703.38958333333 | 43762.479166666664 | 43864.436111111114 | 43892.42083333333 | 43951.477418981478 | 44285.48332175926 | 44319.732638888891 | 44375.450694444444 | 44441.448611111111 | 44491.401388888888
Taux Comp Bi-214 à 609 keV | 1.8514 | 1.08 | 0 | 1.6939 | 2.16 | 1.6971 | 1.3886 | 1.5429 | 0 | 1.8514 | 1.2343 | 1.3886 | 2.7771 | 3.0857 | 2.4321 | 0 | 1.5429 | 2.16 | 1.8514 | 0.89379 | 2.16 | 1.5429 | 16.9714 | 7.56 | 7.7143 | 2.9314 | 2.4686 | 0
Chart: 
| 42366.474305555559 | 42426.53125 | 42489.461261574077 | 42552.5 | 42608.445196759261 | 42671.436793981484 | 42741.476388888892 | 42790.5 | 42857.618750000001 | 42913.404675925929 | 42969.469027777777 | 43053.459918981483 | 43368.474317129629 | 43402.401006944441 | 43461.488078703704 | 43524.442199074074 | 43584.496967592589 | 43644.461111111108 | 43703.38958333333 | 43762.479166666664 | 43864.436111111114 | 43892.42083333333 | 43951.477418981478 | 44285.48332175926 | 44319.732638888891 | 44375.450694444444 | 44441.448611111111 | 44491.401388888888
Taux Comp gamme 30 keV - max | 298.6971 | 643.2171 | 282.4971 | 290.5835 | 296.5371 | 376.7657 | 310.1143 | 295.3029 | 325.3886 | 279.4114 | 306.4114 | 290.6743 | 283.7314 | 278.3314 | 295.9071 | 285.7371 | 288.2057 | 292.0629 | 284.6571 | 279.7572 | 284.9657 | 290.2114 | 554.3486 | 1000.08 | 0 | 811.2343 | 893.4686 | 878.1943
Total : 300  900 d-1
x 3
Electronics issues !
19/03/2021
Partage

## Slide 17: Conclusion
4 HPGe detectors at Laboratoire Souterrain de Modane
Well-type detectors
  High efficiency
  Ultra low background
  Wide versatility
Coaxial detectors
  Need of standard sources for the efficiency calibration of the new coaxial detector (LYS)
  Detection efficiency transfer by simulation (Gespecor) in case of “exotic” sample
Partage
  Many electronics issues:
    Computers
    Acquisition electronics : analogical, digital (DSA1000)…
    Reactivity of LSM staff !
  Increase of the background due to radon contribution
    Waiting for the tent…

## Slide 18
Thanks to LSM staff for their technical support
Thanks to my IRSN colleagues for their ever more challenging samples : Olivier Masson, Céline Duffa, Frédérique Eyrolle, Olivier RadakovicH…
Thank you for your attention