## Slide 1: MedAustron Synchrotron RF
MedAustron
C.Schmitzer, M.Paoluzzi, S.Kouzue, W.Pirkl, U.Dorda,and M.Benedikt
CERN
M.-E.Angoletta, A.Butterworth, J.Molendijk, J.Sanchez-Quesada, M.Jaussi, E.Rigutto, R.Claret

## Slide 2: MedAustron Project
Medical Accelerator for Proton/Carbon ion Radiotherapy & non-clinical research in Wiener Neustadt (1h-Vienna, Austria)
Based on CERN PIMMS
Funded by city of Wr.Neustadt, Country of Lower Austria and the Austrian state
EBG for planning, construction and maintenance of machines and medical operation

## Slide 3: CERN Collaboration
The MedAustron accelerator is based on the CERN PIMMS study and its technical implementation by CNAO/IT.
For the design and construction of the accelerator a collaboration between CERN and MedAustron was established
Cavity design based on CERN PS-Booster (PSB) hardware
  Developed and Prototype tested at CERN
  All parts of Cavity produced by CERN Workshop
  Vacuum chamber is completely assembled by CERN WS
LLRF modified version of new PSB LLRF
  Modified by CERN (BE-RF) for MedAustron application
  Amplifier based on CERN design by M.Paoluzzi
Senior RF Workpackage management by W.Pirkl, since mid 2012 by M.Paoluzzi

## Slide 4: MedAustron Project
MA accelerator complex
10 m
Notes: Split Fodo
Circumference 77 m
Norm Emitt (1 sig) 0.52 – 0.75 pi mm mrad
Bmax= 1.5 T

## Slide 5: MA & LEIR parameters
 | MedAustron | LEIR
Tunes [h/v] | 1.789 / 1.7392 | 1.82/2.72
Injection Energy [MeV/u] | 7 | 4.2
Rigidity Inj/Extr [Tm] | p: 1.137 / 4.881 C: 3.254 / 6.346 | Pb: 1.138 / 4.8
Rev. Period Inj/Extr [ns] | 2170 / 307 | 2760 / 704
Extraction [MeV/u] | p: 60-250 (800) C: 120-400 | Pb: 72.2
Rep Rate [s] | 2 | 3.6
Intensities | 1010 p+ , 4x108 C6+ | 109 Pb54+
Notes: MA extraction resonance: Qx = 1.7392, Qy = 1.779
f = 0.46 - 3.25 MHz
Bmax 1.5 T

## Slide 6: MedAustron ITS at CERN
ECR Ion Source
Spectrometer dipole
Quadrupole triplet
Solenoid
RFQ
BD tank: 4 Slits,
2 Wire Scanners,
1 Faraday Cup
BD tank:
2 Wire Scanners
BD tank:
2 Wire Scanners,
1 Profile Grid,
1 Faraday Cup
Steerer
BD tank:
2 Wire Scanners,
1 Profile Grid,
1 Faraday Cup
BD tank: 4 Slits,
2 Wire Scanners,
1 Profile Grid
Steerer
Fast Deflector
Phase Probe
Current Transformer
Buncher
Doublet-Steerer
Courtesy of W.Prikl
Notes: RFQ succesfully tested
RFQ 75% at 400 microA with 400 keV, (500 us pulses)
Separate buncher from RFQ

## Slide 7: Existing Cavities
 | CNAO | SIEMENS (3) | MedAustron | LEIR
Origin | Saclay/CNAO | Siemens/GSI | MedAustron | CERN
Frequency [MHz] | 0.4 - 4 (0.3-10) * | 1 – 7 h=2 | 0.46 – 3.25 MHz | 0.35 - 5
Length [m] | 1.50 | 1.70 | 2 * 0.7 | 0.60
Voltage [kV] | 4 (8) * | 2.5 | 2x 2.5 | 2.0 ( 4 @ F > 0.7 MHz)
Core material | A-metal | Ferrite | B- metal | B- metal
Bandwidth | Tuned | Tuned | Wideband | Wideband
Core cooling | Water (plates) | Air | Water (plates) + Air | Direct water
Amplifier kind | Tubes | Transistor | Transistor | Tubes
Amplifier location | Ring | Remote | Remote | Ring
Remarks | * achievable with amplifier upgrade | ESR- Cavity from GSI |  | Data for single cavity (2 in-stalled)
A-metal meaning “Mue-Metall”, Cobalt-alloy, amorphous structure,
B-metal meaning “FINEMET, METGLASS” microcristalline structure.

## Slide 8: RF Voltage Requirements
Chart: Frequency [MHz]
| 0.0 | 0.4 | 0.4 | 0.45 | 0.45 | 0.8 | 1.0 | 1.34 | 1.34 | 3.5 | 3.5
 | 0 | 0 | 0.3 | 0.3 | 1.08 | 1.92 | 2.4 | 2.4 | 2.4 | 2.4 | 0

 | 0 | 0 | 0.3 | 0.3 | 1.08 | 2.3 | 3 | 3 | 5 | 5 | 0
MATAC 2011 - W.Prikl
Injection
Trapping
Acceleration
Extraction
Notes: Injection efficiency 0.2 -0.25

## Slide 9: MedAustron Synchrotron Cycle
Horizontal multiturn injection (coasting beam)
RF trapping
  Adiabatic capture VRF≈ 17 - 170 V
  Filling factor K=0.7
  h=1
Acceleration
  @ low energy Voltage and phase ramped for constant bucket area, then ramp to Vpeak = 3kV
  3 T/s (~0.5s)
RF jump to unstable phase for ~ 0.1 ms, back to stable phase for TSync/2
RF-off debunching
Prepare Extraction
Ramp Betatron core for 3rd order slow resonant extraction
RF channeling for extr. smoothing
U.Dorda
Notes: No beam loading
adiabatic factor alpha=0.1
  Rev. Freq. for X MeV/u is identical p/C
  Energy gain per turn for protons x2 higher
To reduce space charge effects (Protons) bunch has large longitudinal extension at low energies
As the phase-space area of the bunch is conserved during the adiabatic voltage ramp, the trapping doubles
the momentum spread (bunch length is halved).
PrepareExtr:
Adjust tune with lattice quads
  Adjust chromaticity with lattice sextupoles
  Turn on RF for channeling
  Ramp resonance sextupole

## Slide 10: Tuning the Momentum Spread
Notes: b for Energy spread
D for homogenous distribution
Delta p/p = 4 E-3 needed for extraction (no matter the energy)
Spill 1-10s
0.1 – 10 s (non clinical)
PrepareExtr:
Adjust tune with lattice quads
  Adjust chromaticity with lattice sextupoles
  Turn on RF for channeling
  Ramp resonance sextupole

## Slide 11: MedAustron Synchrotron Cavity

## Slide 12: Cavity Cell Protoype

## Slide 13: MedAustron Synchrotron Cavity
Joint Venture with Booster C04 upgrade Project
Finemet loaded co-axial resonator
2 blocks à 6 cells (82-121 mm)
2 Finemet FT-3L rings per cell (per gap)
12x 1 kW transistor Amp.
Water cooled Cu disc 815 W 74 °C @ 1.5 l/min Tin= 11.5 °C
Air cooling (cell:58 m3/h)
f = 0.46 - 3.25 MHz
2x 2.5 kV
Notes: common core size 330 x 192 x 25 mm
Finemet (nanocrystalline soft magnetic)
µ’=4800
µ”=5200
@1 MHz

## Slide 14: MedAustron Synchrotron Cavity
2x 6-gap vacuum chamber (à 70 cm)
Ceramic Al2O3 97-98%
Mo/Mn+Ni Metalisation for brazing
Titanium doping for R=10 MΩ
< 2.0x10‐10 mbar l s‐1
Flexible RF contacts
Constructed at Cern WS
Notes: Collar : DILVER P1® (FeNi29Co17)
16 (30) mm gap width

## Slide 15: CERN 1kW RF Amplifier ( M.Paoluzzi)
-1 dB Bandwidth : 0.5 – 5 MHz
Voltage Gain: 58 dB ± 1.5 dB
Output power at 1dB compression : >1 kW
Harmonic distortion : < -15 dBc @ Pout=1 kW
< -18 dBc @ Pout=500 W
Gain linearity: ± 1 dB
DC to RF efficiency : >35 %
Input impedance : 50 Ω, VSWR 1.1:1 max
Load impedance : 50 Ω, no damage when operated in O.C or S.C.
Notes: 60db + voltage factor 1000

## Slide 16: CERN 1kW RF Amplifier ( M.Paoluzzi)
15 W preamplifier
9x 120 W amplifier modules (MRF151G) mounted in push-pull configuration and operated in class B
Signal combiner
Directional coupler (-46dB)
DB Elettronica Telec.
Used @ LEIR, AD,..
Notes: Mosfets push pull
Use N Channel to amplify positive part
P channel to amplify negative part of signal

## Slide 17: MedAustron Cavity Amplifier Chain
Amplifiers are not next to cavity (35 m)
  Accessible during operation
  Resulting in a gain ripple due to cable length
  LLRF Servo Loop has to compensate
LLRF
4W driver
Splitter
1kW
Cell
Gap Voltage
Forward Voltage
Reflected Voltage
0 dB
frequency
1kW
Cell
1kW
…
:
Notes: 2mW for full power
16 times splitter
4 W driver

## Slide 18: MedAustron LLRF - CERN
Developed in Collaboration with Cern BE-RF
Mod. version of PSB LLRF
3-4 DSP boards (VME64x)
VITA57.1 standard FPGA Mezzanine Card (FMC)
Each motherboard with Virtex5 FPGA
Daughtercards 16-bit ADCs and DACs
M.-E.Angoletta et al.
See other talks by J.Molendijk, P.Leinonen & J.Sanchez Quesada
Notes: VME Switched Serial (VXS) enhancement
high-speed P0 connector
BBQ ?

## Slide 19: MedAustron LLRF - Modifications
No synchronisation Loop
Fixed fINJ and fEXTR
Fixed h=1
Additional hardware timings (Extraction, EndDSP,..)
# TPUs, LPUs
12 cells seen as one voltage gap (redundancy)
Tune Kicker output (Ramp up/down timing)
B-field information
reference function inputs
beam diagnostics inputs
cavity return signal
timing inputs
cavity
fREV output
Notes: NO acceleration after the synchrotron
Synchrotron Energy output has to be correct ( Fextr fixed)

## Slide 20: sRF - Slow Control
-Manages Cavity and sRF equipment
-Controls main Power supply (400 VAC & 230 VAC)
-Status and control of 12x PCO and 12x Amplifiers
-Watercooling flowmeters (RFCC & Amps)
-Electrovalves (RFCC & Amps)
-Ventilators
-Thermal switches

## Slide 21: Cavity Test Stand - CTS
Cavity is currently being assembled
Dedicated test stand at Cern (June-Oct.)
  MedAustron control system integrated
  CERN/MA LLRF system
  MA Slow Control
  Beam Interlock System integrated
  Full amplifier chain integrated in Slow Control
  MA Cavity with vacuum chamber

## Slide 22
Thank you for your attention!

## Slide 24
Cavity Voltage step control for calibration
Cav.to gap voltage conversion factor (operator action if cell fails)
ADC to cavity voltage conversion (freq. dependent calibration curve)

## Slide 25: RFQ output beam energy spectrum
H3 beam transmission through the RFQ 75% at 400 microA with 400 keV

## Slide 26: Betatron core
Notes: A BETATRON CORE FOR OPTIMISED SLOW EXTRACTION IN A
PROTON/ION MEDICAL SYNCHROTRON
P.Knauss